4.  System components

The inspection unit consists of:

- metal frame with access doors to the components and subassemblies inside the equipment;

- 2D X-ray detector

- two laser beam scanning systems

- 1 conveyors  to move the object through the laser and  x-ray examination area;

- logic unit;

- Uninterruptible Power Supply (UPS);

- cooling system for the X-ray generator

- image display monitor, mouse and keyboard

The system components and the main elements of control, operation, and signaling are shown in the figures below:

- - the main body of the equipment

- - display

3.  -  keyboard

4. -  cooler

5. -  mobile covers

4

1

3

5

2

Fig.7  - Left side view of the DensityX Laser & X-Ray Scanner equipment

Fig. 8 - Front view of the DensityX -  Laser & X-Ray Scanner equipment

1. - the control panel of the equipment

2. - samples entry door

3. - UPS

4. - PC

5. - display

5

4

3

2

1

Fig. 9 - Rear  view of the DensityX -  Laser & X-Ray Scanner equipment

- – fixed back cover

- - display

3.  -  fans

4. -  cooler

3

4

2

1

3

Fig. 10 - Right-side view of the DensityX -  Laser & X-Ray Scanner equipment

3

3

2

1

3

5. Characteristics

| Power supply | single-phase 230 V +/- 10% - 50/ Hz , 820 VA, 800W |
| --- | --- |
| Conveyor speed | 32 mm/sec |
| Conveyor direction | double direction |
| RX generator | high-frequency - single block (75 kHz) |
| Voltage to RX tube | max.150 kV |
| Current to RX tube | 2.5mA |
| Display | 1 LCD touch-screen monitor |
| Escaping dosage at 10 cm | ≤ 1 µSv/h (0.1 mrem/h) |
| Operating temperature | 0° ÷ +40° C |
| Humidity | 95% without condensate |
| Storage conditions | -20° ÷ +60° C, in closed and dry place |
| Weight | 450 kg |
| Overall dimensions | View Fig.6 |

Fig.12  - Front view of the equipment DensityX -  Laser & X-Ray Scanner with all doors open

600

–––

1940

–––

840

- - mobile covers

- - display

3.   - keyboard

4.   - cooler

–––

2200

–––

840

–––

900

–––

6.   Appearance and dimensions of DensityX -  Laser & X-Ray Scanner

An image representing projections in 3 coordinate axes of the DensityX -  Laser & X-Ray Scanner equipment is presented in Fig.13.

Fig.13  - Appearance and dimensions of DensityX -  Laser & X-Ray Scanner equipment

1700

–––

1040

–––

560

–––

230

–––

1500

Fig. 11 -  Identification plate of the DensityX -  Laser & X-Ray Scanner equipment

150

–––

280

–––

1100

–––

900

–––

7. General Specifications  AION 160kV X-Ray generator

X--ray tube:				Fixed anode glass tube

Tube voltage:				30 - 160 kV

Tube voltage accuracy:			 ±3% (standard)

Tube current:				1 - 7 mA

Tube current accuracy:			±3% (standard)

Duty cycle:				Continuous

Tube max power

water cooling:				500 W (@ TH2O < 22°C; flow = 10 lt/m)

Focal spot:				1,2 mm according to EN 12543

Beam angle:

Total filtration:				0,852 mm Al

Half Value Layer (HVL) @70kV:		1,434 mm Al

Leakage radiations::

@70kV:				< 2,3 mR/h @1m

@90kV:				< 7,85 mR/h @1m

Rise time (kV e mA):			< 100 ms

7.1. Supply Specifications

Mains supply voltage:			230 Vac ±10%

Mains supply frequency:			50/60 Hz

Mains supply power (max):		1 kVA

9. The front side of the equipment

The image presented in Fig.13 show a picture of the front side of the equipment, the elements located here being the following:

| 1. | Control and command panel |
| --- | --- |
| 2. | Lower part of the front side of the equipment |
| 3. | LCD touch-screen display |
| 4. | Ethernet connection socket |

1

2

3

4

Fig.14  - The front side of DensityX -  Laser & X-Ray Scanner equipment

:

| 1. | Sliding door | By lifting it, access to the sample loading part of the equipment is possible, allowing the start of the testing procedure |
| --- | --- | --- |
| 2. | "X-Ray ON" indicator | Light indicator indicating the presence of X-ray emission |
| 3. | "POWER" switch | Used to power the equipment |
| 4. | "Main Power ON" indicator | Shows that the equipment is powered |
| 5. | "X-Ray Power" switch | Used to power the x-ray source |
| 6. | "X-Ray Power ON" indicator | Shows that the X-ray source is powered |
| 7. | "Emergency STOP" switch | Stops the power supply of the X-ray generator |

10.  Control and command panel of the DensityX equipment

On the control panel located on the front of the equipment, there are placed control elements and indicators; their functional description is made below:

Fig.15 - Control Panel of DensityX -  Laser & X-Ray Scanner equipment

2

1

6

5

1

4

3

7